## Slide 1: THz SIS devices and waveguides
AETHER Sub-Task 4.2 & 4.3
KOSMA: University of Cologne
Karl Jacobs, Netty Honingh

## Slide 2: THz SIS devices_ junctions
1.1 THz < fRF < 1.5 THz
4.56 mV < fRF < 6.21 mV
Adding a minimal bias voltage range of ~ 0.5mV:
5.6 mV < 2DE1 + 2DE2 < 7.2 mV
Nb-AlN-Nb(Ti)N
  .

## Slide 3: Simulated Pumped IV‘s

## Slide 4: Nb-AlN-Nb
Reproducible
Quality > 10 , for Jc~30kA/cm2

## Slide 5: THz SIS devices
No progress on junction top electrode other than Niobium (resources)
E-beam lithography required for tuning
  Can be based on heritage of HEB fabrication
  Details will be determined when SIS junction with sufficient Vgap are there.

## Slide 6: THz Waveguides
Last years PR meeting
Stamped waveguides in CuTe up to 5 THz
2. waveguide block with IF output
10 mm
17 mm
1. feedhorn
3. clamp
IF Signal
RF Signal
3. device
This years PR meeting

## Slide 7: In house Tooling
48 µm
16µm
Waveguide stamp
Channel stamp
1 cm
Stamping tool
4.7 THz waveguide 48 µm x 24 µm

## Slide 8: Waveguide wall roughness
roughness Ra < 0.2 µm, RMS = 0.3
8
required accuracy | width (µm) |  | length (µm) |  | depth (µm) | 
Waveguide | 24 | +/- 2 | 48 | +/- 3 | 22 | +/- 2
Substrate Channel | 14 | +2/ -0 | 105 | +0/ -5 | 4 | +2/- 0
4.7 THz

## Slide 9: 4.7 THz HEB Devices
4.7 THz waveguide HEB device
9
1250 µm
IF output and DC connection
RF input section

## Slide 10: Device Assembly
Device aligned to waveguide by windows in beamlead
All electrical contacts via beamleads

## Slide 11
Device aligned to waveguide
11
48 µm
24 µm
HEB microbridge
waveguide
window
LFA device

## Slide 12: Waveguides etched in Si
Less roughness
cheaper series- production
Tests under way at 1.9 THz
50 µm
50 µm
Platzhalter SEM Bild

## Slide 13: Measurement & design
Simulation in CST_microwave Studio
Materials properties at THz frequencies modelled analitically, based on measured DC cryo-properties

## Slide 14: Beam verification
backup mixer
courtesy
C. Risacher, MPIfR
Horns fabricated by RPG
optics path not evacuated ca. 35 cm @ 4775 GHz
waist size and beam shape good to – 12 dB testing limit
flight mixer
contours 1.5 dB
red = -3 dB
green = -12 dB

## Slide 15: GREAT with 1.9 and 4.7 THz Rx on SOFIA

## Slide 16: In-Flight Trec GREAT H-Channel
Trec SSB (K)

## Slide 17: [O I] and [C II] line emission
courtesy Helmut Wiesemeyer
GREAT consortium

## Slide 18: Summary
SIS devices
  AlN barrier: stable process
  New KOSMA sputter machine very probable
  E-beam processing heritage of HEB development
THz waveguide blocks
  Metal waveguides OK up to 5 THz
  Si waveguides ready---tests delayed
Assembly
  Within required tolerance using beamleads
  Cryo thermal contact OK, cryo cycling OK

## Slide 19: Horizon 2020—KOSMA proposal
ALMA second generation instrumentation
=> small FPA‘s with balanced side band separating mixers
Common European effort = our JRA
  Mixers , LO‘s , IF (LNA), …
Frequency range 460GHz - 1 THz(?)
Common prototype + tests + evaluation (result of JRA)
What can each participant contribute ?
  e.g. KOSMA Fabrication of devices, SOI+Beamleads
How can we make the cooperation work ?